{"text_format":"1","workbook":"xlsx","n_sheets":13,"sheets":["Sygehus og kildeoplysninger","Thorax liggende","Thorax stående, siddende","Kurver, Thorax","Oversigt o. Abd.","Kurver, Overs. o. Abdomen","Bækken","Kurver, Bækken","Bækken og hofter","Kurver, Bækken og hofter","Andre","Lister","Data_kurver"],"not_rendered":{"charts":8,"images":0,"macros":0,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[{"sheet":"Kurver, Thorax","index":3,"reason":"not_a_worksheet"},{"sheet":"Kurver, Overs. o. Abdomen","index":5,"reason":"not_a_worksheet"},{"sheet":"Kurver, Bækken","index":7,"reason":"not_a_worksheet"},{"sheet":"Kurver, Bækken og hofter","index":9,"reason":"not_a_worksheet"}],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Sygehus og kildeoplysninger","index":0,"n_rows":22,"n_cols":2,"rows_kept":22,"cut":false,"hidden":false,"merged":[[1,0,1,1],[14,0,20,1],[0,0,0,1],[12,0,13,1],[21,0,40,1]]}
["Konventionelle røntgenundersøgelser \npå børn 0-17 år"]
["Sygehus- og kildeoplysninger"]
["Sygehus"]
["Afdeling"]
["SOR-ID eller AFD-ID "]
["Rum-nr."]
["SIS-registrerings nummer STG-xxxxxxxx"]
["Kilde-fabrikat"," "]
["Model"]
["Yderligere oplysninger"]
["Oplysninger om dosisreduktion"]
["Fokusstørrelser"]
["Kommentarer"]
[]
[]
[]
[]
[]
[]
[]
[]
["Dette dokument benyttes i forbindelse med indsamling og indsendelse af patientdoser på børn ved konventionelle røntgenundersøgelser. Der skal udfyldes et exceldokument pr. røntgenrum. Vær opmærksom på der findes 2 indtastningsark til børn: konventionelle røntgenundersøgelser og CT-skanning.\n\nPå fane for hver undersøgelsestype udfyldes oplysninger om protokol øverst og oplysninger om de enkelte patientundersøgelser nederst. For patientundersøgelserne er de celler, der som minimum skal udfyldes, markeret med blåt. \n\nInformation om hvordan patientdoser på børn måles kan findes i vejledning på SIS' hjemmeside www.sis.dk: Patientdosimetri ved røntgenundersøgelser af børn, 2022.\n\nNB: Projektmappen er beskyttet, således at der ikke kan slettes, indsættes eller omdøbes ark. Ligeledes er de enkelte ark beskyttet, således at der ikke kan ændres i oplysningerne i toppen af arket.\n\nSundhedsstyrelsen, Strålebeskyttelse, April 2022 "]

{"sheet":"Thorax liggende","index":1,"n_rows":45,"n_cols":13,"rows_kept":45,"cut":false,"hidden":false,"merged":[[12,11,12,12],[44,0,44,9],[45,0,51,9],[7,5,7,9],[8,5,8,9],[9,5,9,9],[10,5,10,9],[11,0,11,9],[43,0,43,9],[6,5,6,9],[0,0,1,9],[2,0,2,9],[3,5,3,9],[4,5,4,9],[5,5,5,9]]}
["Thorax, liggende - børn"]
[]
["Protokoloplysninger"]
["Rum:"]
["Undersøgelse:",null,null,null,null,"RU thorax"]
["SKS:",null,null,null,null,"UXRC00"]
["Raster:"]
["Støjreduktion:"]
["FFA cm:"]
["Supplerende filter (Cu, Al):"]
["Fokus størrelse:"]
[]
["Højde [cm]","Vægt [kg]","Køn [K/M]","Alder [år]","Antal optagelser","KAP [Gy·cm²]","Spænding [kV]","Dosisautomatik","mAs","Kommentarer herunder raster eller kompression",null,"Referencedoser"]
[null,null,null,null,null,null,null,null,null,null,null,"Alder","KAP [Gy·cm²]"]
[null,null,null,null,null,null,null,null,null,null,null,"[<5 kg]",1.2999999999999999E-2]
[null,null,null,null,null,null,null,null,null,null,null,"[5-<15 kg]",0.03]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null,null,null,null,null,null,null,null,null," "]
[]
["Kommentarer"]

{"sheet":"Thorax stående, siddende","index":2,"n_rows":45,"n_cols":13,"rows_kept":45,"cut":false,"hidden":false,"merged":[[12,11,12,12],[45,0,51,9],[4,5,4,9],[0,0,1,9],[3,5,3,9],[5,5,5,9],[9,5,9,9],[10,5,10,9],[11,0,11,9],[43,0,43,9],[44,0,44,9],[2,0,2,9],[6,5,6,9],[7,5,7,9],[8,5,8,9]]}
["Thorax, siddende eller stående - børn"]
[]
["Protokoloplysninger"]
["Rum:"]
["Undersøgelse:",null,null,null,null,"RU thorax"]
["SKS:",null,null,null,null,"UXRC00"]
["Raster:"]
["Støjreduktion:"]
["FFA cm:"]
["Supplerende filter (Cu, Al):"]
["Fokus størrelse:"]
[]
["Højde [cm]","Vægt [kg]","Køn [K/M]","Alder [år]","Antal optagelser","KAP [Gy·cm²]","Spænding [kV]","Dosisautomatik","mAs","Kommentarer herunder raster eller kompression",null,"Referencedoser"]
[null,null,null,null,null,null,null,null,null,null,null,"Alder","KAP [Gy·cm²]"]
[null,null,null,null,null,null,null,null,null,null,null,"[15-<30 kg]",2.8000000000000001E-2]
[null,null,null,null,null,null,null,null,null,null,null,"[30-<50 kg]",0.05]
[null,null,null,null,null,null,null,null,null,null,null,"[50-<70 kg]",9.7000000000000003E-2]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null,null,null,null,null,null,null,null,null," "]
[]
["Kommentarer"]

{"sheet":"Oversigt o. Abd.","index":4,"n_rows":45,"n_cols":14,"rows_kept":45,"cut":false,"hidden":false,"merged":[[12,12,12,13],[45,0,51,10],[4,5,4,10],[0,0,1,10],[3,5,3,10],[5,5,5,10],[9,5,9,10],[10,5,10,10],[11,0,11,10],[43,0,43,10],[44,0,44,10],[2,0,2,10],[6,5,6,10],[7,5,7,10],[8,5,8,10]]}
["Oversigt over abdomen - børn"]
[]
["Protokoloplysninger"]
["Rum:"]
["Undersøgelse:",null,null,null,null,"Oversigt over abdomen"]
["SKS:",null,null,null,null,"UXRD00"]
["Raster:"]
["Støjreduktion:"]
["FFA cm:"]
["Supplerende filter (Cu, Al):"]
["Fokus størrelse:"]
[]
["Højde [cm]","Vægt [kg]","Køn [K/M]","Alder [år]","Antal optagelser","KAP [Gy·cm²]","Spænding [kV]","Dosisautomatik","mAs","Projektioner","Kommentarer herunder raster eller kompression",null,"Referencedoser"]
[null,null,null,null,null,null,null,null,null,null,null,null,"Alder","KAP [Gy·cm²]"]
[null,null,null,null,null,null,null,null,null,null,null,null,"[5-<15 kg]",7.2999999999999995E-2]
[null,null,null,null,null,null,null,null,null,null,null,null,"[15-<30kg]",0.23699999999999999]
[null,null,null,null,null,null,null,null,null,null,null,null,"[30-<50 kg]",0.53400000000000003]
[null,null,null,null,null,null,null,null,null,null,null,null,"[50-<70 kg]",1.54]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
["Kommentarer"]

{"sheet":"Bækken","index":6,"n_rows":45,"n_cols":13,"rows_kept":45,"cut":false,"hidden":false,"merged":[[12,11,12,12],[45,0,51,9],[0,0,1,9],[3,5,3,9],[4,5,4,9],[5,5,5,9],[9,5,9,9],[10,5,10,9],[11,0,11,9],[43,0,43,9],[44,0,44,9],[2,0,2,9],[6,5,6,9],[7,5,7,9],[8,5,8,9]]}
["Bækken - børn"]
[]
["Protokoloplysninger"]
["Rum:"]
["Undersøgelse:",null,null,null,null,"Bækken,  f.eks. dysplasi eller metastaser"]
["SKS:",null,null,null,null,"UXRG10"]
["Raster:"]
["Støjreduktion:"]
["FFA cm:"]
["Supplerende filter (Cu, Al):"]
["Fokus størrelse:"]
[]
["Højde [cm]","Vægt [kg]","Køn [K/M]","Alder [år]","Antal optagelser","KAP [Gy·cm²]","Spænding [kV]","Dosisautomatik","mAs","Kommentarer herunder raster eller kompression",null,"Referencedoser"]
[null,null,null,null,null,null,null,null,null,null,null,"Alder","KAP [Gy·cm²]"]
[null,null,null,null,null,null,null,null,null,null,null,"[5-<15 kg]",4.1000000000000002E-2]
[null,null,null,null,null,null,null,null,null,null,null,"[15-<30kg]",0.13]
[null,null,null,null,null,null,null,null,null,null,null,"[30-<50 kg]",0.33]
[null,null,null,null,null,null,null,null,null,null,null,"[50-<70 kg]",0.64600000000000002]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
["Kommentarer"]

{"sheet":"Bækken og hofter","index":8,"n_rows":45,"n_cols":14,"rows_kept":45,"cut":false,"hidden":false,"merged":[[12,12,12,13],[45,0,51,10],[0,0,1,10],[3,5,3,10],[4,5,4,10],[5,5,5,10],[9,5,9,10],[10,5,10,10],[11,0,11,10],[43,0,43,10],[44,0,44,10],[2,0,2,10],[6,5,6,10],[7,5,7,10],[8,5,8,10]]}
["Bækken og hofter - børn"]
[]
["Protokoloplysninger"]
["Rum:"]
["Undersøgelse:",null,null,null,null,"Bækken og hofter, f.eks. Smerter, patologisk gang eller kontrol"]
["SKS:",null,null,null,null,"UXRG10,  UXRG15, "]
["Raster:"]
["Støjreduktion:"]
["FFA cm:"]
["Supplerende filter (Cu, Al):"]
["Fokus størrelse:"]
[]
["Højde [cm]","Vægt [kg]","Køn [K/M]","Alder [år]","Antal optagelser","KAP [Gy·cm²]","Spænding [kV]","Dosisautomatik","mAs","Projektioner","Kommentarer herunder raster eller kompression",null,"Referencedoser"]
[null,null,null,null,null,null,null,null,null,null,null,null,"Alder","KAP [Gy·cm²]"]
[null,null,null,null,null,null,null,null,null,null,null,null,"[5-<15 kg]",4.8000000000000001E-2]
[null,null,null,null,null,null,null,null,null,null,null,null,"[15-<30kg]",8.7999999999999995E-2]
[]
[]
[]
[]
[null,null,null,null,null,null,null," "]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
["Kommentarer"]

{"sheet":"Andre","index":10,"n_rows":45,"n_cols":11,"rows_kept":45,"cut":false,"hidden":false,"merged":[[45,0,51,10],[0,0,1,10],[3,5,3,10],[4,5,4,10],[5,5,5,10],[9,5,9,10],[10,5,10,10],[11,0,11,10],[43,0,43,10],[44,0,44,10],[2,0,2,10],[6,5,6,10],[7,5,7,10],[8,5,8,10]]}
["Indtast protokol navn - børn"]
[]
["Protokoloplysninger"]
["Rum:"]
["Undersøgelse:"]
["SKS:"]
["Raster:"]
["Støjreduktion:"]
["FFA cm:"]
["Supplerende filter (Cu, Al):"]
["Fokus størrelse:"]
[]
["Højde [cm]","Vægt [kg]","Køn [K/M]","Alder [år]","Antal optagelser","KAP [Gy·cm²]","Spænding [kV]","Dosisautomatik","mAs","Projektioner","Kommentarer herunder raster eller kompression"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
["Kommentarer"]

{"sheet":"Lister","index":11,"n_rows":3,"n_cols":9,"rows_kept":3,"cut":false,"hidden":true,"merged":[]}
[null,null,null,null,null,null,"År ",null,"Mdr"]
["Stå/sidde","AP/PA","1. gang","Aksial","Ja"]
["Liggende","AP/PA + lat.","Kontrol","Lauenstein","Nej"]

{"sheet":"Data_kurver","index":12,"n_rows":83,"n_cols":21,"rows_kept":83,"cut":false,"hidden":true,"merged":[]}
["Thorax",null,null,null,null,null,"Oversigt over abdomen",null,null,null,null,null,"Bækken",null,null,null,null,"Bækken og hofter"]
[]
["[5-<15 kg]",null,0.03,null,null,null,"[5-<15 kg]",null,7.2999999999999995E-2,null,null,null,"[5-<15 kg]",null,4.1000000000000002E-2,null,null,"[5-<15 kg]",null,4.8000000000000001E-2]
["[15-<30kg]",null,2.8000000000000001E-2,null,null,null,"[15-<30kg]",null,0.23699999999999999,null,null,null,"[15-<30kg]",null,0.13,null,null,"[15-<30kg]",null,8.7999999999999995E-2]
["[30-<50 kg]",null,0.05,null,null,null,"[30-<50 kg]",null,0.53400000000000003,null,null,null,"[30-<50 kg]",null,0.33]
["[50-<70 kg]",null,9.7000000000000003E-2,null,null,null,"[50-<70 kg]",null,1.54,null,null,null,"[50-<70 kg]",null,0.64600000000000002]
[]
[]
[null,"k","DQ2","DQ3",null,null,null,"k","DQ2","DQ3",null,null,null,"k","DQ2","DQ3",null,null,"k","DQ2","DQ3"]
["Samlet",null,null,null,null,null,"Samlet",5.6000000000000001E-2,2.8400000000000002E-2,6.08E-2,null,null,"Samlet",6.5000000000000002E-2,1.4200000000000001E-2,2.4299999999999999E-2,null,"Samlet",0.14399999999999999,5.1999999999999998E-3,1.0500000000000001E-2]
["Liggende",0.14729999999999999,5.3745485368193683E-3,7.8227057187457878E-3,null,null,"Et billede",4.7E-2,1.78E-2,3.49E-2]
["Siddende/stående",3.5900000000000001E-2,7.4790000000000004E-3,1.2751999999999999E-2,null,null,"To billeder",6.7000000000000004E-2,3.7100000000000001E-2,7.0800000000000002E-2]
[]
[]
[]
[null,"Liggende",null,"Siddende",null,null,null,"Samlet",null,"1 Projektion",null,null,null,"Samlet",null,null,null,null,"Samlet"]
["w","Q2","Q3","Q2","Q3",null,"w","Q2 - samlet","Q3 - samlet","Q2 - 1 pr","Q3 - 1 projektion",null,"w","Q2","Q3",null,null,"w","Q2 - samlet","  Bækken - alle"]
[5,{"$f":"$C$11*EXP($B$11*$A18)","v":1.1225349593564984E-2},{"$f":"$D$11*EXP($B$11*$A18)","v":1.6338601439530132E-2},{"$f":"$C$12*EXP($B$12*$A18)","v":8.949512783520689E-3},{"$f":"$D$12*EXP($B$12*$A18)","v":1.5259284264668514E-2},null,5,{"$f":"$I$10*EXP($H$10*$G18)","v":3.7576886670383211E-2},{"$f":"$J$10*EXP($H$10*$G18)","v":8.0446292590116167E-2},{"$f":"$I$11*EXP($H$11*$G18)","v":2.2515376083445473E-2},{"$f":"$J$11*EXP($H$11*$G18)","v":4.4145316028777921E-2},null,5,{"$f":"$O$10*EXP($N$10*M18)","v":1.9653235172926671E-2},{"$f":"$P$10*EXP($N$10*M18)","v":3.3631944697332258E-2},null,null,5,{"$f":"Data_kurver!$T$10*EXP(Data_kurver!$S$10*$R18)","v":1.0683052695348215E-2},{"$f":"Data_kurver!$U$10*EXP(Data_kurver!$S$10*$R18)","v":2.157154871176082E-2}]
[6,1.300682965174829E-2,1.8931562344713323E-2,9.2766370401932332E-3,1.5817044462701442E-2,null,6,{"$f":"$I$10*EXP($H$10*$G19)","v":3.9741228304630428E-2},{"$f":"$J$10*EXP($H$10*$G19)","v":8.5079812708504579E-2},{"$f":"$I$11*EXP($H$11*$G19)","v":2.3598861215764898E-2},{"$f":"$J$11*EXP($H$11*$G19)","v":4.6269677327539047E-2},null,6,{"$f":"$O$10*EXP($N$10*M19)","v":2.0973127273133527E-2},{"$f":"$P$10*EXP($N$10*M19)","v":3.5890633291348215E-2},null,null,6,{"$f":"Data_kurver!$T$10*EXP(Data_kurver!$S$10*R19)","v":1.23376877883919E-2},{"$f":"Data_kurver!$U$10*EXP(Data_kurver!$S$10*R19)","v":2.4912638803483645E-2}]
[7,1.5071033305419807E-2,2.1936030090349914E-2,9.6157184035700199E-3,1.6395192015286119E-2,null,7,4.2030231003824997E-2,8.9980212853259134E-2,2.4734485829459467E-2,4.8496267160007614E-2,null,7,2.2381662039082678E-2,3.8301013207725984E-2,null,null,7,{"$f":"Data_kurver!$T$10*EXP(Data_kurver!$S$10*R20)","v":1.4248599562755693E-2},{"$f":"Data_kurver!$U$10*EXP(Data_kurver!$S$10*R20)","v":2.8771210655564386E-2}]
[8,1.7462829219305028E-2,2.5417311438065752E-2,9.9671939320403989E-3,1.6994472124799993E-2,null,8,4.4451074956560019E-2,9.5162864695734109E-2,2.5924758981125333E-2,5.0830004968610905E-2,null,8,2.388479262572419E-2,4.0873271887682937E-2,null,null,8,{"$f":"Data_kurver!$T$10*EXP(Data_kurver!$S$10*R21)","v":1.6455481203761581E-2},{"$f":"Data_kurver!$U$10*EXP(Data_kurver!$S$10*R21)","v":3.3227413969133965E-2}]
[9,2.0234206783482318E-2,2.9451077431911141E-2,1.033151665943329E-2,1.7615657232396484E-2,null,9,4.7011353913660361E-2,0.10064402527994894,2.7172310468202929E-2,5.327604692922934E-2,null,9,2.5488871996086581E-2,4.3618280951049568E-2,null,null,9,{"$f":"Data_kurver!$T$10*EXP(Data_kurver!$S$10*R22)","v":1.9004173740354662E-2},{"$f":"Data_kurver!$U$10*EXP(Data_kurver!$S$10*R22)","v":3.8373812360331536E-2}]
[10,2.3445406183329538E-2,3.4125008225749306E-2,1.0709156178954436E-2,1.825954801364179E-2,null,10,4.9719099008409276E-2,0.10644088801800296,2.8479896639269014E-2,5.5839797343285875E-2,null,10,2.7200679771997329E-2,4.6547642145037676E-2,null,null,10,{"$f":"Data_kurver!$T$10*EXP(Data_kurver!$S$10*R23)","v":2.194761824838207E-2},{"$f":"Data_kurver!$U$10*EXP(Data_kurver!$S$10*R23)","v":4.4317306078463799E-2}]
[11,2.7166227813290422E-2,3.9540698947253768E-2,1.1100599248467825E-2,1.8926974410544417E-2,null,11,5.258280394876489E-2,0.11257163662270793,2.9850406483931568E-2,5.8526920578045608E-2,null,11,2.9027450888071512E-2,4.9673736378882935E-2,null,null,11,{"$f":"Data_kurver!$T$10*EXP(Data_kurver!$S$10*R24)","v":2.5346955535027856E-2},{"$f":"Data_kurver!$U$10*EXP(Data_kurver!$S$10*R24)","v":5.1181352522652407E-2}]
[12,3.1477549496597729E-2,4.5815868025421543E-2,1.1506350417901464E-2,1.9618796701307589E-2,null,12,5.5611451660590089E-2,0.11905550214661539,3.1286868015782739E-2,6.134335358150661E-2,null,12,3.0976906170074456E-2,5.3009776051606283E-2,null,null,12,{"$f":"Data_kurver!$T$10*EXP(Data_kurver!$S$10*R25)","v":2.9272796146891263E-2},{"$f":"Data_kurver!$U$10*EXP(Data_kurver!$S$10*R25)","v":5.9108530681222751E-2}]
[13,3.6473084490074748E-2,5.308691598302244E-2,1.1926932679586227E-2,2.033590660918352E-2,null,13,5.881454246547059E-2,0.12591282330635958,3.2792454962512221E-2,6.4295318999532391E-2,null,13,3.3057284966897332E-2,5.6569860894056696E-2,null,null,13,{"$f":"Data_kurver!$T$10*EXP(Data_kurver!$S$10*R26)","v":3.3806687082134051E-2},{"$f":"Data_kurver!$U$10*EXP(Data_kurver!$S$10*R26)","v":6.8263502762001452E-2}]
[14,4.2261418486973248E-2,6.1511890313302717E-2,1.2362888142365983E-2,2.107922845185867E-2,null,14,6.2202123881510266E-2,0.1331651102815431,3.4370493777962433E-2,6.7389338924207245E-2,null,14,3.5277379974064396E-2,6.0369037561251039E-2,null,null,14,{"$f":"Data_kurver!$T$10*EXP(Data_kurver!$S$10*R27)","v":3.9042805673031103E-2},{"$f":"Data_kurver!$U$10*EXP(Data_kurver!$S$10*R27)","v":7.8836434532082042E-2}]
[15,4.8968369895260933E-2,7.1273920887132386E-2,1.2814778730347946E-2,2.1849720332851583E-2,null,15,6.5784822140582996E-2,0.14083511218829037,3.6024470991617784E-2,7.0632249303789932E-2,null,15,3.7646574395953022E-2,6.4423363226877348E-2,null,null,15,{"$f":"Data_kurver!$T$10*EXP(Data_kurver!$S$10*R28)","v":4.5089915824009952E-2},{"$f":"Data_kurver!$U$10*EXP(Data_kurver!$S$10*R28)","v":9.1046945413866251E-2}]
[16,5.6739724695663754E-2,8.2585200564493103E-2,1.3283186907193868E-2,2.2648375376458908E-2,null,16,6.9573875520905534E-2,0.14894688843912168,3.7758040911766211E-2,7.4031215046103424E-2,null,16,4.0174881603792151E-2,6.8749973448742896E-2,null,null,16,{"$f":"Data_kurver!$T$10*EXP(Data_kurver!$S$10*R29)","v":5.2073627239875131E-2},{"$f":"Data_kurver!$U$10*EXP(Data_kurver!$S$10*R29)","v":0.10514867038820941}]
[17,6.5744405325840388E-2,9.5691597535065218E-2,1.376871642688565E-2,2.3476223007841395E-2,null,17,7.3581169599489016E-2,0.15752588421299057,3.9575033699352238E-2,7.7593745848729953E-2,null,17,4.2872987457053278E-2,7.3367154591999606E-2,null,null,17,{"$f":"Data_kurver!$T$10*EXP(Data_kurver!$S$10*R30)","v":6.013900457258186E-2},{"$f":"Data_kurver!$U$10*EXP(Data_kurver!$S$10*R30)","v":0.12143452846386722}]
[18,7.6178142471296104E-2,0.11087799964428295,1.4271993111933091E-2,2.433433027989982E-2,null,18,7.7819274535050331E-2,0.16659901027222043,4.1479463830360136E-2,8.1327712790987008E-2,null,18,4.5752295467104694E-2,7.8294421116242532E-2,null,null,18,{"$f":"Data_kurver!$T$10*EXP(Data_kurver!$S$10*R31)","v":6.945358068338188E-2},{"$f":"Data_kurver!$U$10*EXP(Data_kurver!$S$10*R31)","v":0.14024280714913651}]
[19,8.8267729575100637E-2,0.12847450687206494,1.4793665660026835E-2,2.5223803248651182E-2,null,19,8.2301484498332689E-2,0.17619472737671221,4.3475538965423978E-2,8.5241365724342516E-2,null,19,4.8824974994012775E-2,8.355259805313453E-2,null,null,19,{"$f":"Data_kurver!$T$10*EXP(Data_kurver!$S$10*R32)","v":8.0210836611390654E-2},{"$f":"Data_kurver!$U$10*EXP(Data_kurver!$S$10*R32)","v":0.16196418931146192}]
[20,0.10227595254477118,0.14886360656734055,1.5334406480176222E-2,2.6145788398877812E-2,null,20,8.7041859373521269E-2,0.18634313556021453,4.5567669246261031E-2,8.9343351499691567E-2,null,20,5.2104012680193275E-2,8.9163909023147636E-2,null,null,20,{"$f":"Data_kurver!$T$10*EXP(Data_kurver!$S$10*R33)","v":9.2634220533983416E-2},{"$f":"Data_kurver!$U$10*EXP(Data_kurver!$S$10*R33)","v":0.18704986838592808}]
[21,0.11850730181113711,0.17248848740321129,1.5894912559409835E-2,2.7101474121887174E-2,null,21,9.2055268861563927E-2,0.19707606854870022,4.7760477039468355E-2,9.3642733071766598E-2,null,21,5.5603267338296652E-2,9.5152070163423141E-2]
[22,0.13731459089963699,0.19986267276944494,1.6475906361155822E-2,2.8092092247286943E-2,null,22,9.7357439123741008E-2,0.208427193617023,5.0058807149210335E-2,9.8149009522889921E-2,null,22,5.9337528525310108E-2,0.1015423903637349]
[23,0.15910662537894926,0.2315811829990142,1.707813675646001E-2,2.9118919630749838E-2,null,23,0.10296500211179663,0.22043211719708572,5.2467737521361364E-2,0.10287213705030963,null,23,6.3322579050439129E-2,0.10836187823420215]
[24,0.18435708888343802,0.26833346905696864,1.7702379989242036E-2,3.0183279799814736E-2,null,24,0.10889554773937141,0.23312849656879511,5.4992590462753309E-2,0.10782255096348824,null,24,6.7575261680957538E-2,0.11563935625684985]
[25,0.21361483936087935,0.31091839882540734,1.8349440676833747E-2,3.1286544659845421E-2,null,25,0.11516767905838879,0.24655615798415625,5.7638944400315702E-2,0.11301118873994484,null,25,7.211355032795716E-2,0.12340558260347598]
[26,0.24751583935036664,0.36026162173467624,1.9020152847089432E-2,3.2430136262345824E-2,null,26,0.12180107061347339,0.26075722159504161,6.0412646206089458E-2,0.11844951419059113,null,26,7.6956626012863127E-2,0.13169338113468829]
[27,0.28679697961346862,0.41743568918795348,1.9715381013404719E-2,3.3615528637911077E-2,null,27,0.12881653015744887,0.27577623357651027,6.3319824115344911E-2,0.12414954278795155,null,27,8.2124957935786211E-2,0.14053778012954962]
[28,0.33231209659668398,0.48368336812782808,2.043602128902974E-2,3.4844249696176922E-2,null,28,0.13623606392150528,0.29166030586012398,6.6366901266345743E-2,0.13012386821322847,null,28,8.7640389988346473E-2,0.14997616033217034]
[29,0.38505053188955424,0.5604446545013605,2.118300254211291E-2,3.6117883195216444E-2,null,29,0.14408294564477692,0.30845926391557871,6.9560609891673439E-2,0.13638569018086535,null,29,9.3526233076613979E-2,0.16004841294096614]
[30,0.44615863709701709,0.649388073803149,2.1957287592964045E-2,3.7438070782922515E-2,null,30,0.15238178957986409,0.32622580304421606,7.2908006192466929E-2,0.14294884360208404,null,30,9.9807363644367986E-2,0.17079710820831984]
[31,0.5169646915677143,0.75244694906219378,2.2759874455080129E-2,3.8806514113007326E-2,null,31,0.16115862770330394,0.34501565367467885,7.6416485928440414E-2,0.14982782915182982,null,31,0.10651032881308208,0.18226767536323199]
[32,0.59900777460369503,0.8718614246134585,2.359179762153317E-2,4.0224977038346169E-2,null,32,0.17044099137318514,0.36488775617921321,8.0093800758124209E-2,0.15703784530665929,null,32,0.11366345858300822,0.19450859461740139]
[33,0.69407121973372243,1.0102271591059098,2.4454129398378471E-2,4.1695287884492879E-2,null,33,0.18025799769005232,0.38590444575898525,8.3948075365430794E-2,0.16459482192435587,null,33,0.12129698556957613,0.20757160206624645]
[34,0.8042213782306491,1.1705517461650112,2.5347981286801807E-2,4.3219341806297176E-2,null,34,0.1906404408319988,0.4081316479783636,8.7987825410385662E-2,0.17251545543946403,null,34,0.12944317278117234,0.22151190835087942]
[35,0.93185253445797145,1.3563200891001879,2.6274505415787131E-2,4.4799103230661511E-2,null,35,0.20162088865044839,0.43163908556152325,9.2221976343683412E-2,0.18081724575250285,null,35,0.13813644997835037,0.23638843200520515]
[36,1.0797389493502163,1.5715701506778263,2.7234896027151432E-2,4.643660838858605E-2,null,36,0.21323378482963143,0.45650049710005597,9.6659883126637924E-2,0.18951853489436313,null,36,0.14741355919079219,0.25226404847438377]
[37,{"$f":"$C$11*EXP($B$11*$A50)","v":1.2510951632728435},{"$f":"$D$11*EXP($B$11*$A50)","v":1.8209807244984968},{"$f":"$C$12*EXP($B$12*$A50)","v":2.8230391014860821E-2},{"$f":"$D$12*EXP($B$12*$A50)","v":4.8133967939765362E-2},null,37,0.22551555693020917,0.48279386835763083,0.10131135090009645,0.19863854755131272,null,37,0.15731371000704727,0.26920585585713019]
[38,1.4496458690377532,2.1099731358252627,2.9262273520612007E-2,4.9893369693119971E-2,null,38,0.23850473065596109,0.51060167689726876,0.10618665664798306,0.20819743354014658,null,38,0.16787874529337837,0.28728545849500664]
[39,1.6797068738726511,2.4448290824881633,3.0331873587735789E-2,5.1717081426769188E-2,null,39,0.25224205070196881,0.54001115079858109,0.11129657190333558,0.21821631232732652,null,39,0.17915331804212403,0.3065792696072967]
[40,1.9462789101781339,2.8328271773195275,3.1440569875554283E-2,5.3607453811080119E-2,null,40,0.26677060856337509,0.57111454227652125,0.11665238654700244,0.22871731969047107,null,40,0.19118508009702403,0.32716883425054111]
[41,2.255156334194639,3.282401160081494,3.2589791436401734E-2,5.5566923438560623E-2,null,41,0.2821359777056317,0.60400941705994382,0.1222659337515808,0.23972365662529044,null,41,0.20402488355315387,0.34914117396772104]
[42,2.6130530753132057,3.8033232178671237,3.3781019557599212E-2,5.7598015964501287E-2,null,42,0.29838635652024015,0.63879896043769724,0.12814961612570694,0.25125964060602091,null,42,0.2177269956826822,0.37258915458374481]
[43,3.027748573733462,4.4069163987280566,3.5015789670757601E-2,5.9703349362414879E-2,null,43,0.31557271951441268,0.67559230093226375,0.13431643311646257,0.26335075931261487,null,43,0.23234932829483246,0.39761187870172027]
[44,3.5082568786499673,5.106300735668043,3.6295693330869759E-2,6.1885637298469201E-2,null,44,0.33374897720890617,0.71450485261624985,0.14078000973043997,0.27602372694339072,null,44,0.24795368249943597,0.42431510455889387]
[45,4.0650225825779867,5.9166784308887017,3.7622380267742825E-2,6.4147692629262801E-2,null,45,0.35297214524560228,0.75565867714551471,0.14755462663692226,0.28930654323756105,null,45,0.26460600990856592,0.45281169301254581]
[46,4.7101478507548373,6.8556642992089136,3.8997560512414813E-2,6.6492431027452018E-2,null,46,0.37330252323529167,0.79918286664456806,0.15465525171968936,0.3032285553380426,null,46,0.28237669138021915,0.48322208454502291]
[47,5.4576554804526198,7.943668653357606,4.0423006601296503E-2,6.8922874739902787E-2,null,47,0.39480388390667703,0.84521394864528032,0.1620975731471607,0.31782052263123084,null,47,0.30134083448215093,0.51567480830396251]
[48,6.3237937082040974,9.2043409537450085,4.1900555860879374E-2,7.1442156483210811E-2,null,48,0.41754367314991958,0.89389631434912353,0.16989803403393855,0.33311468470699185,null,48,0.32157859093309066,0.55030702532916209]
[49,7.3273893903990457,10.665083866127578,4.343211277595474E-2,7.4053523481611816E-2,null,49,0.44159322158223169,0.94538267155632694,0.17807386877033532,0.34914483258902823,null,49,0.34317549536300035,0.58726510826203582]
[50,8.4902572341785358,12.357648901006362,4.501965144439643E-2,7.6760341652486061E-2,null,50,0.46702796829915638,0.99983452368270098,0.18664314110015084,0.3659463833929924,null,50,0.36622282682414148,0.62670525998779136]
[51,9.8376739738947148,14.318826581904068,4.666521812167107E-2,7.9566099944852184E-2,null,51,0.4939276975134031,1.057422676366722,0.1956247840308277,0.38355645857729698,null,51,0.39081799458087196,0.66879417382501316]
[52,11.398927800096173,16.591246144397704,4.8370933858355961E-2,8.247441483644273E-2,null,52,0.52237678882353267,1.1183277732560135,0.20503864166416486,0.40201396595951427,null,52,0.41706494980870995,0.71370973805293314]
[53,13.207954983729159,19.224302148604366,5.0138997234063731E-2,8.5489034995157198E-2,null,53,0.55246448189753705,1.1827408626538822,0.2149055130400081,0.42135968567956644,null,53,0.44507462494270533,0.7616417877540661]
[54,15.304077533568051,22.275228146478145,5.1971687191298195E-2,8.8613846110901784E-2,null,54,0.5842851564015823,1.2508639968033874,0.22524719808978738,0.44163636029963932,null,54,0.47496540253201791,0.81279290714986152]
[55,17.732857921002218,25.810340741740475,5.3871365972893631E-2,9.1852875904043257E-2,null,55,0.61793862805199717,1.3229108656887825,0.23608654580143115,0.46288878923988469,null,55,0.50686361558231352,0.86737928582043788]
[56,20.547089450946988,29.906481084014128,5.5840482166823996E-2,9.5210299316932673E-2,null,56,0.65353046171917151,1.3991074673424517,0.24744750470207219,0.48516392775855732,null,56,0.54090408150066616,0.92563163242719615]
[57,23.807943806124857,34.652685130269326,5.7881573862306503E-2,9.8690443895190855E-2,null,57,0.69117230256551399,1.4796928167599734,0.25935517577008338,0.50851099069527583,null,57,0.57723067189967869,0.98779614980015418]
[58,27.586300708369922,40.152118979302301,5.9997271921267803E-2,0.10229779536569153,null,58,0.73098222625618781,1.5649196956470499,0.27183586789334702,0.53298156120661866,null,58,0.61599692066909306,1.0541355755111943]
[59,31.964288599204249,46.524321346740052,6.2190303369389167E-2,0.10603700341843168,null,59,0.77308510934117569,1.6550554453501225,0.28491715599628681,0.55862970473429274,null,59,0.65736667288489614,1.1249302923311955]
[60,37.037069828765389,53.907801924239571,6.4463494911102134E-2,0.10991288769974253,null,60,0.81761302097049204,1.7503828054579547,0.2986279419640897,0.58551208845768155,null,60,0.70151477729852874,1.2004795132643835]
[61,42.91490915693312,62.463052102245719,6.6819776573064621E-2,0.11393044402456479,null,61,0.86470563717128568,1.8512008007047243,0.31299851849872456,0.6136881064946903,null,61,0.7486278253329921,1.28110254616843]
[62,49.725570528731467,72.376033499030626,6.9262185480814142E-2,0.11809485081579649,null,62,0.91451067898634242,1.9578256789566766,0.32806063604783897,0.64322001112750449,null,62,0.7989049397092155,1.3671401433052066]
[63,57.617094221633153,83.86222012459848,7.1793869773466107E-2,0.12241147577901319,null,63,0.96718437584835137,2.0705919032246394,0.3438475729544106,0.6741730503431983,null,63,0.85255861603579819,1.4589559415260489]
[64,66.761014730367165,97.171282042152896,7.4418092661502622E-2,0.12688588282116345,null,64,1.0228919556434566,2.1898532008141607,0.36039420898214036,0.70661561199307288,null,64,0.9098156209191004,1.5569379991784604]
[65,77.356089334939838,112.59251233376325,7.7138236632882662E-2,0.13152383922215799,null,65,1.0818081630013323,2.3159836729042604,0.3777371023790343,0.74061937488923013,null,65,0.97091795038952577,1.661500436229963]
[66,89.632618398073973,130.46111533373943,7.9957807812894521E-2,0.13633132306859619,null,66,1.1441178074375675,2.4493789680353553,0.39591457064944086,0.77625946717221828,null,66,1.036123852694786,1.7730851845410771]
[67,103.85745130042397,151.16549281420959,8.2880440483370399E-2,0.14131453095921101,null,67,1.2100163430677815,2.590457523187363,0.41496677521300124,0.81361463229964859,null,67,1.105708919782959,1.8921638556849227]
[68,120.33978682532494,175.15568649943276,8.5909901767088445E-2,0.14647988599196574,null,68,1.2797104817119307,2.7396618763410343,0.43493581013755683,0.85276740302251308,null,68,1.1799672520884814,2.0192397342077535]
[69,139.43789407342726,202.95316041997978,8.9050096483399496E-2,0.15183404604309536,null,69,1.3534188413120003,2.8974600546397751,0.45586579514206355,0.89380428373359655,null,69,1.2592127015440411,2.1548499047549434]
[70,161.56689999670735,235.16213574140232,9.2305072181338044E-2,0.15738391234876623,null,70,{"$f":"$I$10*EXP($H$10*$G83)","v":1.4313726316970599},{"$f":"$J$10*EXP($H$10*$G83)","v":3.0643470425063812},{"$f":"$I$11*EXP($H$11*$G83)","v":0.47780297307499447},{"$f":"$J$11*EXP($H$11*$G83)","v":0.9368159415908599},null,70,{"$f":"$O$10*EXP($N$10*M83)","v":1.3437801980719217},{"$f":"$P$10*EXP($N$10*M83)","v":2.2995675220526546}]